Masson Trichrome Staining Kit

Instruction For Use

[Product Name]

Masson Trichrome Staining Kit

[Specification]

4×10mL/box,, 4×20mL/box, 4×100mL/box

Type: A; B

[Intended Use]

This kit is used for histological staining of connective tissue, muscle and collagen fibers in the tissue sections.

[Principle of Detection]

Masson Trichrome Staining usually refers to staining the nucleus and selectively showing collagen fibers and muscle fibers. This staining principle is related to the size of the anion dye molecule and penetration of the tissue. The size of the molecule is reflected by the molecular weight. Small molecular weights can easily penetrate tissues with dense structures and low permeability, while large molecular weights can only penetrate tissues with loose structure and high permeability. However, the molecular weight of hematoxylin is very large, so after Masson staining, muscle fibers show red, and collagen fibers show green or blue, which is mainly used to distinguish collagen fibers from muscle fibers.

[Main Components]

Hematoxylin solution (Solution A) is mainly composed of hematoxylin;

Ponceau acid magenta staining solution (Solution B) is mainly composed of ponceau and acid magenta;

Phosphomolybdic acid solution (Solution C) is mainly composed of phosphomolybdic acid;

(Type A) Aniline blue solution (Solution D) is mainly composed of aniline blue;

(Type B) Brilliant green solution (Solution D) is mainly composed of brilliant green.

[Storage and validity]

Store at 2~8℃ avoid freezing, valid for 12 months. During the shelf life, it can be transported within 7 days at -20℃-2℃/8℃-30℃. After each use, it should be immediately put back in the refrigerator at 2~8℃ and run out within six months.

[Applicable Instrument]

Optical microscope (100× ~ 400×)

[Sample Requirements]

The tissue is preserved by freezing method, which can be different according to different fixation and preservation conditions. The tissue sections cut by freezing microtome with a thickness of 4 ~ 30μm were spread on sticky slides. Dry them at room temperature for 1 hour, and then stain them.

[Test method]

- Rinse the sections with distilled water for 1min

- Stain with Solution A

- Rinse with tap water for 2 min

- Add Solution B to completely cover the tissue for 5-10min

- Pour Solution B, simple and quick washing (filter paper can absorb excess liquid on the slide), add Solution C to rinse

- Pour Solution C, do not wash with water (filter paper can absorb excess liquid on the slide)

- Add Solution D for counterstaining for 3-5 min, pour the staining solution on the slide (simple, quick wash for several times, and there is no excess solution on the slide)

- Dehydrate with alcohol (fast, gradient alcohol is not easy to soak for a long time),  transparentize with xylene, and sealed with neutral balata

[Positive judgment value/Reference Range]

| Collagen fiber, mucus, cartilage | Blue/Green |
| --- | --- |
| Muscle fiber, cytoplasm, erythrocyte | Red |
| Cell nucleus | Blue |

[Results Interpretation]

It depends on tissue types and purpose of the staining experiment. Any clinical explanation for staining or its absence must be supplemented by the qualified pathologist in combination with path morphology.

[Limitations of detection Methods]

Excessive or insufficient fixation will affect staining quality. The specific staining time of Masson trichrome must be tested by user experiments. The best staining time setting should use a negative control.

[Product Performance]

1.  Consistency

Six frozen sections from the different sample source were taken, diluent is the blank control. The test results are: collagen fibers show blue; muscle fiber, cytoplasm and erythrocyteglycogen show red; the cell nucleus show blue brown. The staining location should be accurate, and the blank control is not chromogenic.

2. In-batch repeatability

Three frozen sections from the same tissue source were taken and stained with reagent of same batch. The test results are: collagen fibers show blue; muscle fiber, cytoplasm and erythrocyteglycogen show red; the cell nucleus show blue brown. The staining location should be accurate

3. Inter-batch repeatability

Three frozen sections from the same tissue source were taken and stained with reagent of same batch. three different batches of reagents respectively. The test results showed that there was no significant difference in staining intensity and location of tissue slices from the same tissue source with different batches of reagents.

[Cautions]

- The sections should be as clean as possible when deparaffinized. The differentiation time of acid alcohol should be determined according to the thickness of the slice, tissue type and the old and new.

- Phosphomolybdic acid will differentiate collagen fibers stained red into colorless or light red.

- After staining, the aniline blue solution is treated with weak acid solution, the purpose is to remove the blue color in the protoplasm, and make the staining show bright and clear.

- 4. This product should be used and judged by professionals.

5. Read the instruction manual carefully before use, use it within the validity period, and take personal hygiene protection.

6. After use, the waste should be disposed according to the requirements of the hospital or environmental protection department.

7. See the outer packaging for the date of manufacture, batch code and expiry date.

[Symbols]

| Symbol | Used for | Symbol | Used for |
| --- | --- | --- | --- |
|  | Use-by date |  | Consult instructions for use |
|  | Batch code |  | In vitro diagnostic medical device |
|  | Temperature limit |  | Manufacturer |
|  | Avoid overexposure to the sun |  | Date of manufacture |
|  | Place upward |  |  |
| [Basic Information] BASIC INFORMATION |  |  |  |  |

Henan Celnovte Biotechnology , Ltd.

Address: N0.1 Cuizhu Street, Bldg 109, Hi-tech District, Zhengzhou, Henan, China. 450000

Tel: +86(371)-56596939

Email: service@celnovte-bio.com